## Slide 1: Healing Humanity One Spark at a Time: Biomedical Applications of a Laser-Induced Plasma
Presented at the University of Waterloo
Wednesday, May 5, 2-17
Steven J. Rehse
Department of Physics
Windsor, Ontario, Canada

## Slide 2: Outline
Introduction of the Method. Laser-induced breakdown spectroscopy (LIBS)
Advantages of LIBS over other analytic methods
Biomedical Applications of LIBS
  A new paradigm for rapid pathogen identification
  A real time assay for nutritional zinc deficiency
Concluding Thoughts

## Slide 3: Laser-Induced Plasmas
What is a “laser-induced plasma?”
Incident Laser Pulse
Laser-induced plasma
Rare-earth metal target
Can be done with ns, ps, or fs lasers
Threshold irradiance:

## Slide 4
absorption of laser energy
pulsed
laser
laser interaction with the target
initiated by absorption of energy by the target from a pulsed radiation field.
pulse durations are on the order of nanoseconds, but can be performed with pico- and femto-second laser pulses.
LIBS primer

## Slide 5
melting
sublimation
fragmentation
crater
vapor
atomization
absorbed energy is rapidly converted into heating, resulting in vaporization of the sample (ablation) when the temperature reaches the boiling point of the material.
removal of particulate matter from the surface leads to the formation of a vapor above the surface.
removal of samples mass (ablation)
LIBS primer

## Slide 6
electrical breakdown and plasma formation
absorption of the laser radiation by the vapor
continuum emission
shock wave
bremmstrahlung
plasma formation (breakdown)
The laser pulse continues to illuminate the vapor plume.
The vapor condenses into sub-micrometer droplets that lead to absorption and scattering of the laser beam, inducing strong heating, ionization, and plasma formation.
LIBS primer

## Slide 7
The dynamic evolution of the plasma plume is then characterized by a fast expansion and subsequent cooling.
Approximately 1 microsecond after the ablation pulse, spectroscopically narrow atomic/ionic emissions may be identified in the spectrum.
crater
debris
expansion and element specific emission (atomic or ionic)
spontaneous emission
as atoms/ions decay to ground state
LIBS primer

## Slide 9: Laser-Induced Plasmas
What is a “laser-induced plasma?”
Incident Laser Pulse
Laser-induced plasma
Rare-earth metal target
Can be done with ns, ps, or fs lasers
Threshold irradiance:
We can do spectroscopy on that!

## Slide 10: What It looks Like

## Slide 11: The Goal of LIBS Plasma Creation
to create an optically thin plasma which is in thermodynamic equilibrium (or LTE) and whose elemental composition is the same as that of the target/sample
  if achieved, atomic emission spectral line intensities can be related to relative concentrations of elements (sometimes absolute concentrations)
  typically these conditions are only met approximately.
steel LIBS spectrum

## Slide 12
When we do a time-resolved spectroscopy of the plasma, we call it:
“Laser-induced breakdown spectroscopy”
or
LIBS
1 ns
10 ns
100 ns
1 μs
10 μs
100 μs
LIBS
laser pulse
ICCD Observation Window
(10 μs)
Gate delay
Elapsed Time After LIBS Pulse Incident on Target
LIBS emission intensity

## Slide 13: History
1960
Maiman: first ruby laser
1962
Brech, Cross: Birth of LIBS: detection of spectrum from ruby laser induced plasma
1964
Runger et al. First direct spectro-chemical analysis by LIBS
1965
Zel’dovich, Raizer: First theoretical model for laser breakdown of a gas

## Slide 14: History
1962-64: First spectrochemical analysis of laser-induced plasma
LIBS coined at LANL
>2000 peer-reviewed papers have been published according to the Web of Science
Why the interest?
Applications & Advantages!
“laser-induced breakdown” and/or “laser-induced plasma spectroscopy” @ Web of Science (Thomson Reuters)

## Slide 15: Applications of LIBS
No matter what your application is, you will be doing one of two things:
Attempting to quantify the amount/concentration of some element by analyzing peak intensities
Attempting to identify a target based on its unique elemental composition by analyzing the presence intensity of all/many lines

## Slide 16: Outline
Introduction of the Method. Laser-induced breakdown spectroscopy (LIBS)
Advantages of LIBS over other analytic methods
Biomedical Applications of LIBS
  A new paradigm for rapid pathogen identification
  A real time assay for nutritional zinc deficiency
Concluding Thoughts

## Slide 17: Keeping Track of the Elemental Inventory (underlined elements reported in literature)
1 H |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  |  | 2 He
3 Li | 4 Be |  |  |  |  |  |  |  |  |  |  | 5 B | 6 C | 7 N | 8 O | 9 F | 10 Ne
11 Na | 12 Mg |  |  |  |  |  |  |  |  |  |  | 13 Al | 14 Si | 15 P | 16 S | 17 Cl | 18 Ar
19 K | 20 Ca | 21 Sc | 22 Ti | 23 V | 24 Cr | 25 Mn | 26 Fe | 27 Co | 28 Ni | 29 Cu | 30 Zn | 31 Ga | 32 Ge | 33 As | 34 Se | 35 Br | 36 Kr
37 Rb | 38 Sr | 39 Y | 40 Zr | 41 Nb | 42 Mo | 43 Tc | 44 Ru | 45 Rh | 46 Pd | 47 Ag | 48 Cd | 49 In | 50 Sn | 51 Sb | 52 Te | 53 I | 54 Xe
55 Cs | 56 Ba | 57 La | 72 Hf | 73 Ta | 74 W | 75 Re | 76 Os | 77 Ir | 78 Pt | 79 Au | 80 Hg | 81 Tl | 82 Pb | 83 Bi | 84 Po | 85 At | 86 Rn
87 Fr | 88 Ra | 89 Ac | 104 Rf | 105 Db | 106 Sg | 107 Bh | 108 Hs | 109 Mt | 110 Uun | 111 Uuu | 112 Uub |  | 114 Uuq |  | 116 Uuh |  | 
 |  |  | 58 Ce | 59 Pr | 60 Nd | 61 Pm | 62 Sm | 63 Eu | 64 Gd | 65 Tb | 66 Dy | 67 Ho | 68 Er | 69 Tm | 70 Yb | 71 Lu | 
 |  |  | 90 Th | 91 Pa | 92 U | 93 Np | 94 Pu | 95 Am | 96 Cm | 97 Bk | 98 Cf | 99 Es | 100 Fm | 101 Md | 102 No | 103 Lr | 
(■ Solids ■ Liquids ■ Gases ■ Artificially prepared)
www.arl.army.mil/wmrd/LIBS

## Slide 18: Advantages of LIBS - spatial resolution
Laser allows point sampling (1-100 micron)
Elemental “surface maps” can then be created
Courtesy of Ben Smith, Javier Laserna
Courtesy of F.C. Alvira et al.

## Slide 19: Advantages of LIBS - depth profiling
Because laser only removes g to ng of material, ablation crater only microns deep
Subsequent shots thus sample progressively deeper layers
Courtesy of Ben Smith
Courtesy of Ben Smith

## Slide 20: Advantages of LIBS – sensitivity & speed
Concentrations of 1-100 ppm usually detectable in seconds using a standard LIBS apparatus

## Slide 21: Advantages of LIBS - portability / standoff
Apparatus is compact, low weight; can be made man-portable
All optical technique, so can be done at a distance “stand-off”

## Slide 22
First responder CBRNE prototypes have been built…
Head’s-up display
Hand-held probe contains laser, joystick for control, and focus optics
Microplasma/ LIBS Event
Backpack contains broadband high-resolution spectrometer, laser power supply, computer, and battery
courtesy of Ocean Optics.

## Slide 23
Courtesy of A.J. Miziolek, A. Whitehouse
High-energy remote systems have been built…

## Slide 24
Commercial benchtop systems have been built…
ChemReveal LIBS Desktop Elemental Analyzer – TSO
J200 – Applied Spectra

## Slide 25
Hand-held systems have been built…
mPulse – Oxford Instruments
ChemLite- TSI, Inc
LIBZ – SciApps, Inc
EOS500 - Bruker
NanoLIBS – B&WTek

## Slide 26: And a system has gone to Mars…

## Slide 27: And a system has gone to Mars…

## Slide 28: And a system has gone to Mars…
SELFIE ON MARS

## Slide 29
LIBS Table 1
1064 nm LIBS laser
355 nm OPO pump
OPO
Argon chamber
Echelle spectrometer/camera
Timing control
Photodiode for observing pulse timing
LIBS Table 2
1064 nm LIBS laser

## Slide 31
IR Nd:YAG laser
Echelle
spectrometer
sample chamber
(w/ purge gas)
microscope
assembly

## Slide 33: Outline
Introduction of the Method. Laser-induced breakdown spectroscopy (LIBS)
Advantages of LIBS over other analytic methods
Biomedical Applications of LIBS
  A new paradigm for rapid pathogen identification
  A real time assay for nutritional zinc deficiency
Concluding Thoughts

## Slide 35: So why?
“It is well-accepted that the microbiological expertise and cost required to perform these identifications preclude their common use as a screening mechanism to prevent human infection.”1
1Tarr, P.I. 1995. Escherichia coli O157:H7: clinical, diagnostic, and epidemiological aspects of human infection. Clin. Infect. Dis. 20, 1-8.

## Slide 36: Current Method of Bacterial identification

## Slide 37: How unique is “unique”?
We can identify a bacterial species, certainly its genus, with high sensitivity and specificity (confirmed by others).
We can differentiate strains of E. coli (demonstrated by others in MRSA).
Multiple multivariate techniques effective at discriminating spectra.

## Slide 38: Results: We have already demonstrated…
LIBS spectral fingerprint is a sensitive and specific (high rates of true positives, low rates of false positives) test to identify an unknown bacterial specimen or to differentiate between possible identifications
This spectral fingerprint is robust and reliable, and exists through time (multiple tests spanning years on same strains of bacteria)
In addition…
9 publications in Applied Physics Letters, Journal of Applied Physics, Applied Optics, Applied Spectroscopy, Spectrochimica Acta B, and others – confirmed by multiple other groups

## Slide 39: Results: We have already demonstrated…
LIBS spectral fingerprint is:
  growth-medium independent
  independent of state of growth (how “old” the bacteria are)
  independent of whether the bacteria are live or dead (or inactivated by UV light)
  obtainable even when other types of bacteria or contaminants are present (mixed samples)
  obtainable from urine specimens
  capable of strain discrimination
  obtainable from about 500 bacteria
9 publications in Applied Physics Letters, Journal of Applied Physics, Applied Optics, Applied Spectroscopy, Spectrochimica Acta B, and others – confirmed by multiple other groups

## Slide 40: Outline
Introduction of the Method. Laser-induced breakdown spectroscopy (LIBS)
Advantages of LIBS over other analytic methods
Biomedical Applications of LIBS
  A new paradigm for rapid pathogen identification
  A real time assay for nutritional zinc deficiency
Concluding Thoughts

## Slide 41: Motivation for this Project
~25 receptors
Frederickson et al., Nature Neuroscience, 2003
Zn2+
Zinc deficiency is a leading cause of death among toddlers worldwide.
It is also a leading cause of weakened immunity in the elderly.
Copenhagen Consensus (8 Nobel Laureates) says: Zinc supplementation is the number ONE most cost-effective move to improve world health. (2-3 billion people deficient worldwide).

## Slide 42
How do we diagnose and monitor zinc deficiency & remediation in
millions of people ?
Why not fingernail zinc?
Zinc in the fingernails has been shown to represent the overall zinc concentration in the body.
Need a real-time biomedical assay!
Can LIBS do this?

## Slide 43: Preparation of Nails
Nail clippings of the index, middle and ring fingers (both right and left hands) of 5 subjects were taken a total of 6 nail clippings per subject.
Clippings were cleaned with acetone in an ultrasound bath for 10 minutes and allowed to dry for 20-30 minutes.
Clippings are cut into approximately 2 mm by 1 mm fragments to provide a flat target.

## Slide 44: Typical Nail Components
Farren, Shayler, Ennos, The Journal of Experimental Biology, 2004
C247.9
Ca393.4
Ca396.8
Ca422.7
Mg279.6
Mg280.3
Mg285.2
Ca315.7
Ca317.9
Na589.0
Na589.6
Ar
CN
Ar
Zn202.5
Zn206.2
Zn213.8
Wavelength (nm)
Element | Mean (ppm) | ±Std.Dev.
Mg | 570.8 | ±511.5
Al | 837.4 | ±427.2
Si | 974.3 | ±594.1
P | 1035.2 | ±597.8
S | 24810 | ±8776
Cl | 1770.2 | ±1121.9
K | 831.9 | ±666.1
Ca | 2311.6 | ±1448.3
Fe | 128.6 | ±69.6
Ni | 5.1 | ±2.1
Cu | 20.8 | ±20.8
Zn | 151.7 | ±74.8
Olabanji et al. Nuc. Instrum. Methods Section B: Beam Interactions with Materials and Atoms, 2005

## Slide 45: Observed Zinc Lines
ZnII 202.547
ZnII 206.2
ZnI 213.855

## Slide 46: A Comparison of LIBS-Predicted Zinc Content vs. Actual Content
“Determination of the Zinc Concentration in Human Fingernails Using Laser-Induced Breakdown Spectroscopy,” Vlora A. Riberdy, Christopher J. Frederickson, and Steven J. Rehse, Applied Spectroscopy 71(4), 567-582, 2017
SIDMS: speciated isotope dilution mass spectrometry.

## Slide 47: A Comparison of LIBS-Predicted Zinc Content vs. Actual Content
Other things we’ve studied:
Smoothness of surface
Hydration state of nail
Sampling location along growth direction
Sampling depth (# of laser shots)
Layered structure
Laser wavelength
“Determination of the Zinc Concentration in Human Fingernails Using Laser-Induced Breakdown Spectroscopy,” Vlora A. Riberdy, Christopher J. Frederickson, and Steven J. Rehse, Applied Spectroscopy 71(4), 567-582, 2017

## Slide 48: Conclusions
LIBS provides an accurate, fast, spatially resolved, remote spectrochemical analysis of almost any type of target (solid, liquid, gas, powder)
High degree of versatility and robustness suggests its adoption in many different interdisciplinary fields
Experiments utilizing LIBS involve an exciting mixture of physics, laser science, and analytic chemistry (at a minimum)

## Slide 49: Funding and Acknowledgements
We gratefully acknowledge funding for this project provided by:
A Natural Sciences and Engineering Research Council of Canada Discovery grant and a Research Tools and Instruments grant
A Canada Foundation for Innovation Leaders Opportunity Fund grant
An Ontario Research Fund Small Infrastructure Funds grant
University of Windsor Outstanding Scholars program
University of Windsor Faculty of Science

## Slide 50
All Credit to the Students!
Khadijia Sheikh
Russell Putnam
Derek Gillies
Dylan Malenfant
Anthony Piazza
Vlora Riberdy
Allie Paulick
Dan Trojand
Siddharth Doshi

## Slide 51
All Credit to the Students!
Courtney Jones
Chris Heath
Beau Greaves
Paul Dubovan
Allie Paulick
Erica Rustico

## Slide 52
Thank you!